| Date | Document No. |
| --- | --- |
| 08-04-2024 | IS 14697:2021 Amendment 2 draft |

| Sl. No. | Name of the Organization | Clause/ Sub-clause | Paragraph/ Figure/Table | Type of Comment (General/ Technical/ Editorial | Comments | Proposed Change |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Schneider Electric India Private Limited | 1.1 | Para 1 | Technical | Frequency range is 45 to 55Hz. Since the frequency variation is only for +/- 5 | %. Should we reduce the range with respect 5%?. To be discussed BMVyas comment : This is generic comments and not relevant to reason (i.e. inclusion of class 0.1s) for which amendement is being done. This is not recommended because the frequency range of 45 to 55 Hz is mentioned in all other standards like 13779, 15884. |
| 2 | Schneider Electric India Private Limited | Point xii) | Table 13 | Technical | Error limit for EMHF test can be defined as 1% instead of 0.5% for 0.1S similar to EFT considering the measurement uncertainty due to challenges in performing accuracy test in EMI environment. | Error limit of 1% for class 0.1S same as class 0.2S meters BMVyas comment : This is not recommended because the limits we have mentioned in our amendement is in line with IEC 62053-22. Relaxing the limit will make quality meter poorer. |
| 3 | Schneider Electric India Private Limited | Annex J | diagram | Technical | In earth fault condition, the U2 & U3 can be replaced with U2=1.9Un & U3=1.9Un. | Earth fault condition diagram to be retained same as IS 14697:2021 BMVyas comment : This is not recommended because the diagram is not meant for the specific value of voltage, which is already specified in relevant clause. The diagram is indicative and is for explaining the phasor arrangement only. Also the diagram is in line with IEC 62052-11 |