भारतीय मानक

Indian Standard

IS 11196 : 2024

संकुचन कारकों के निर्धारण के लिए

उपकरण — विशिष्टि

( पहला पुनरीक्षण )

Equipment for Determination

of Liquid Limit of Soils by

Cone Penetration

Method ― Specification

( First Revision )

ICS 93.020; 13.080.20

BIS 2024

भारतीय मानक ब्यूरो

BUREAU OF INDIAN STANDARDS

मानक भवन, 9 बहादुर शाह ज़फर मार्ग, नई दिल्ली - 110002

MANAK BHAVAN, 9 BAHADUR SHAH ZAFAR MARG

NEW DELHI - 110002

www.bis.gov.in     www.standardsbis.in

October 2024                                             Price Group X

Soil and Foundation Engineering Sectional Committee, CED 43

FOREWORD

This Indian Standard (First Revision) was adopted by the Bureau of Indian Standards, after the draft finalized by the Soil and Foundation Engineering Sectional Committee had been approved by the Civil Engineering Division Council.

There are a series of standards on methods of testing of soils.  It has been recognized that reliable and inter-comparable test results can be obtained only with the standard testing equipment capable of giving the desired level of accuracy.  With this objective, a series of specifications covering the requirements of equipment used for testing soils have been published to encourage their development and manufacturing in the country.

The equipment covered in this standard is used for determination of liquid limit of soil by cone penetration method as covered in IS 2720 (Part 5) : 1985 ‘Methods of the test for soils :    Part 5 Determination of liquid and plastic limits (second revision)’.

This standard was first published in 1985.  The present revision has been taken up with a view to incorporate the modifications found necessary as a result of experience gained in the use of this standard.  Also, in this revision, the standard has been brought into latest style and format of Indian Standards, and references to Indian Standards, wherever applicable have been updated.  The other major modifications incorporated in this revision of the standard are given below:

- Digital type equipment for determination of liquid limit test by cone penetration method has been included.  Provisions on material and construction of digital type equipment have also been included.

- Provision for providing an arrangement of making soil thread for plastic limit test has also been incorporated considering that such arrangements are being provided by the manufacturer along with the equipment for liquid limit determination.

- BIS certification marking clause has been modified to align with the revised Bureau of Indian Standards Act, 2016.

This standard contributes to the Sustainable Development Goal 9 - Industry, Innovation and Infrastructure: Build resilient infrastructure, promote inclusive and sustainable industrialization and foster innovation.

The composition of the Committee responsible for formulation of the standard is given in Annex B.

For the purpose of deciding whether a particular requirement of this standard is complied with, the final value, observed or calculated, expressing the result of a test or analysis, shall be rounded off in accordance with IS 2 : 2022 ‘Rules for rounding off numerical values (second revision)’.  The number of significant places retained in the rounded off value should be the same as that of the specified value in this standard.

Indian Standard

EQUIPMENT FOR DETERMINATION OF LIQUID LIMIT OF

SOILS BY CONE PENETRATION

METHOD — SPECIFICATION

(First Revision)

1  SCOPE

This standard covers the requirements of the equipment for determination of liquid limit of soils by cone penetration method.

2  REFERENCES

The following standards contain provisions, which through reference in this text, constitute provisions of this standard.  At the time of publication, the editions indicated are valid.  All standards are subject to revision, and parties to agreement based on this standard are encouraged to investigate the possibility of applying the most recent editions of the standards is given in Annex B.

- TYPES

  - Equipment for liquid limit test by cone penetration method can be of the following two types:

- Analogue type; and

- Digital type.

4 DIMENSIONS

Dimensions with tolerances of different components of the analogue type equipment shall be as detailed in Figs. 1 to 10.  General assembly of different components of the digital type equipment shall be as detailed in Fig. 11.  Dimensions with tolerance of cup and cone of digital type equipment shall be as detailed in Fig. 9 and Fig. 10.  Except where tolerances are specially mentioned against the dimensions, all dimensions shall be taken as nominal dimensions and tolerances of medium class as given in IS 2102 (Part 1) shall apply thereon.

FIG.  1 GENERAL ASSEMBLY FOR ANALOGUE TYPE EQUIPMENT

All dimensions in millimetres.

FIG.  2 BASE BODY

All dimensions in millimetres.

FIG. 3 PILLAR

FIG. 4 BRACKET

All dimensions in millimetres.

FIG. 5 PLUNGER

6A PUSH BUTTON PIN

6B PUSH BUTTON SPRINGS

All dimensions in millimetres.

Fig. 6 RELEASE MECHANISM

7A DIAL

7D PINION SHAFT

All dimensions in millimetres.

FIG. 7 DIAL GUAGE

All dimensions in millimetres.

FIG. 8 SPUR RACK

All dimensions in millimetres.

FIG. 9 CONE

All dimensions in millimetres.

FIG. 10 CUP

FIG. 11 GENERAL ASSEMBLY OF DIGITAL CONE PENETROMETER FOR SOIL

4 MATERIALS

The materials of construction for various component of the analogue type and digital type equipment shall be as given in Table 1 and Table 2 respectively.

Table 1 Materials of Construction for Various Components of Analogue

Type Equipment

(Clause 4)

| Sl No. | Components Parts | Material | Special Requirement | Relevant IS Number |
| --- | --- | --- | --- | --- |
| (1) | (2) | (3) | (4) | (5) |
|  | Base body | Cast aluminium | – | IS 617 |
|  | Levelling screws | Mild steel | Nickel/Chrome plated | IS 1875 |
|  | Pillar | Mild steel | Nickel/Chrome plated | IS 1875 |
|  | Bracket | Cast aluminium | – | IS 617 |
|  | Knurled clamp | Mild steel | Nickel/Chrome plated | IS 1875 |
|  | Clamp pin | Mild steel | Nickel/Chrome plated | IS 1875 |
|  | Plunger rod | Mild steel | Nickel/Chrome plated | IS 1875 |
|  | Top cap | Mild steel | Nickel/Chrome plated | IS 1875 |
|  | Clamping screw | Mild steel | Nickel/Chrome plated | IS 1875 |
|  | Bush | Brass | – | IS 4170 |
|  | Weight | Brass | – | IS 4170 |
|  | Push button pin | Mild steel | Nickel/Chrome plated | IS 1875 |
|  | Push button-spring | Spring steel | Nickel/Chrome plated | IS 4454 (Part 1) |
|  | Dial | Aluminium | Anodized | IS 7883 |
|  | Needle pointer | Mild steel | Nickel/Chrome plated | IS 1875 |
|  | Pointer bush | Brass | – | IS 4170 |
|  | Pinion shaft | Mild steel | Nickel/Chrome plated | IS 1875 |
|  | Spur rack | Mild steel | Nickel/Chrome plated | IS 1875 |
|  | Cone | Stainless steel | – | IS 6603 |
|  | Cup | Brass | – | IS 4170 |

Table 2  Materials of Construction of Various Component Parts of Digital

Type Equipment

(Clause 4)

| Sl No. | Components Parts | Material | Special Requirement | Relevant IS Number |
| --- | --- | --- | --- | --- |
| (1) | (2) | (3) | (4) | (5) |
|  | Base body | Mild steel | Painted |  |
|  | Levelling screws | Mild steel | Nickel/Chrome plated | IS 1875 |
|  | Pillar | Mild steel | Nickel/Chrome plated | IS 1875 |
|  | Bracket | Cast aluminium | Painted | IS 617 |
|  | Clamp knob | Mild steel | Nickel/Chrome plated | IS 1875 |
|  | Up down Knob | Mild steel | Nickel/Chrome plated | IS 1875 |
|  | Clamping screw | Mild steel | Nickel/Chrome plated | IS 1875 |
|  | Weight | Brass | – | IS 4170 |
|  | Spur rack | Brass | – | IS 4170 |
|  | Cone | Stainless steel | – | IS 6603 |
|  | Cup | Brass | – | IS 4170 |

5 CONSTRUCTION

5.1 Analogue Type

The base, pillar, bracket, plunger, release mechanism, dial gauge, spur rack, cone and cup shall be according to details given in Figs. 1 to 10.  The mass of the plunger along with cone and needle as given in Fig. 5 and Fig. 9 respectively shall be 80 ± 0.5 g (for this purpose thickness of weight given in Fig. 5E be adjusted). When assembled, the pin fixed at the bottom of the rack shall be in line and concentric with the plunger assembly.

5.2 Digital Type

The equipment shall be power operated and shall have facility for adjusting the height of the cone in relation to soil test specimen in the cup.  The linear variable differential transformer (LVDT) provided for test shall be enough to accommodate the cone.  The equipment shall have electrical indicator with visual display.  Provision shall be made to stop the equipment after 5 ± 0.1 s of the penetration of cone.  The mass of cone along with the siding components and cone shall be 80 ± 0.5 g (for this purpose thickness of weight given in Fig. 5E be adjusted).

5.3 A 300 mm x 300 mm non-absorbent plate and 3 mm thick steel rod to perform plastic limit test shall be provided with the liquid limit test apparatus.

#### 6 MARKING

6.1 The following information shall be clearly and indelibly marked suitably on the equipment:

- Name of manufacturer;

- Date of manufacture; and

- Type of apparatus, analogue or digital.

6.2 BIS Certification Marking

The product conforming to the requirements of this standard may be certified as per the conformity assessment schemes under the provisions of the Bureau of Indian Standards Act, 2016 and the Rules and Regulations framed thereunder, and the product may be marked with the Standard Mark.

Annex A

(Clause 2)

| IS No. | Title |
| --- | --- |
| IS 617 : 2024 | Aluminium and aluminium alloys ingots for remelting and castings for general engineering purposes – specification (fourth revision) |
| IS 1875 : 1992 | Carbon steel billets, blooms, slabs and bars for forgings — Specification (fifth revision). |
| IS 2102 (Part 1) : 1993 | General tolerances: Part 1 tolerances for linear and angular dimensions without individual tolerance indications (third revision) |
| IS 2720 (Part 5) : 1985 | Method of test for soils: Part 5 Determination of liquid and plastic limit (second revision) |
| IS 4170 : 1967 | Specification for brass rods for general engineering purposes. |
| IS 4454 (Part 1) : 2001 | Steel wire for mechanical springs — Specification: Part 1 Cold drawn unalloyed steel wire (third revision). |
| IS 7883 : 1975 | Specification for aluminium-manganese alloy sheet and strip for aircraft purposes (Alloy No. 31000) |
| IS 6603 : 2024 | Stainless steel semi-finished products, bars, wire rods and bright bars – specification (second revision) |

ANNEX B

(Foreword)

COMMITTEE COMPOSITION

Soil and Foundation Engineering Sectional Committee, CED 43

| Organization | Representative(s) |
| --- | --- |
| In Personal Capacity, 473, Vinayak Apartments, BHEL Housing Society, Plot No. C-58/19, Sector 62, Noida, Uttar Pradesh - 201301 | Shri C. Pushpakaran (Chairperson) |
| AFCONS Infrastructure Limited, Mumbai | Dr Sunil Basarkar Dr Lakshmana Rao Mantri (Alternate-I) Shri Budhmal Jain (Alternate-II) |
| AIMIL Limited, New Delhi | Shri Rohitash Barua Smt Aarti Bhargava (Alternate-I) Shri Anil Singh (Alternate-II) |
| Bharat Heavy Electricals Ltd, New Delhi | Shri T. M. S. Rao Shri Vikram S. (Young Professional) |
| CEM Engineers and Consultants Pvt Ltd, Bhubaneswar | Shri Ashok Basa Shri Dilip Basa (Alternate) |
| Cengrs Geotechnica Pvt Ltd, Noida | Shri Sanjay Gupta Shri Ravi Sundaram (Alternate) Shri Sorabh Gupta (Young Professional) |
| Central Board of Irrigation and Power, New Delhi | Director |
| Central Electricity Authority, New Delhi | Shri Baleshwar Thakur Shri Deepak Singh Raghuvansi (Alternate) |
| Central Public Works Department, New Delhi | Shri Nagendra Prasad Shri Amrendra Kumar Jalan (Alternate) |
| Central Soil and Materials Research Station, New Delhi | Dr Manish Gupta Ms Swapna Varma (Alternate) |
| CSIR-Central Building Research Institute, Roorkee | Shri Manojit Samanta Dr S. Ganesh Kumar (Alternate) Shri Kaushik Pandit (Young Professional) |
| CSIR-Central Road Research Institute, New Delhi | Dr Kanwar Singh Dr P. S. Prasad (Alternate) |
| CSIR-Structural Engineering Research Centre, Chennai | Dr P. Kamatchi Smt R Sreekala (Alternate) Dr A. Thirumalaiselvi (Young Professional) |
| D-CAD Technologies, New Delhi | Dr K. G. Bhatia |
| Delhi Development Authority, New Delhi | Shri Arun Kumar Shri Harindar Pal (Alternate) |
| Delhi Technological University, New Delhi | Prof. Ashok Kumar Gupta |
| Engineers India Limited, New Delhi | Shri V. K. Panwar Shri Sampat Raj (Alternate-I) Shri Anil Banoth (Young Professional) |
| Geodynamics Ltd, Vadodara | Dr Ravikiran Vaidya Shri Sujan Kulkarni (Alternate) |
| Geological Survey of India, Kolkata | Dr Timir Baran Ghosal Shri Prashant Tukaram Ilamkar (Alternate) |
| Ground Engineering Limited, New Delhi | Shri Ashok Kumar Jain Shri Neeraj Kumar Jain (Alternate) |
| Hindustan Construction Company Limited, Mumbai | Representative |
| Indian Geotechnical Society, New Delhi | Prof H. N. Ramesh Dr Anil Joseph (Alternate) Prof D. Neelima Satyam (Alternate-II) |
| Indian Institute of Science, Bengaluru | Prof Jyant Kumar Prof G. Madhavi Latha (Alternate) |
| Indian Institute of Technology Delhi, New Delhi | Dr G. V. Ramana Dr J. T. Shahu (Alternate-I) Dr Prashanth Vangla (Young Professional) |
| Indian Institute of Technology Kanpur, Kanpur | Prof Priyanka Ghosh |
| Indian Institute of Technology Madras, Chennai | Prof Subhadeep Banerjee Prof Ramesh K Kandasami (Alternate) |
| Indian Institute of Technology Bombay, Mumbai | Prof Deepankar Choudhury Prof Dasaka Murty (Alternate) |
| Indian Institute of Technology Roorkee, Roorkee | Dr Mahendra Singh Dr Vishwas A. Sawant (Alternate) |
| Indian Road Congress, New Delhi | Secretary General Director (T) (Alternate) |
| Indian Society of Earthquake Technology, Roorkee | Prof B. K. Maheswari Prof Vasant A. Matsagar (Alternate) |
| ITD Cementation India Ltd, Kolkata | Shri Manish Kumar Shri Aminul Islam (Alternate) |
| Jadhavpur University, Kolkata | Prof Sibapriya Mukherjee Prof Ramendu Bikas Sahu (Alternate) |
| Keller Ground Engineering Pvt Ltd, Chennai | Shri V. V. S. Ramadas Shri Madan Kumar Annam (Alternate) |
| L&T GeoStructure Private Limited, Chennai | Shri M. Kumaran Shri A. Vetriselvan (Alternate) |
| Military Engineer Services, Engineer-in-Chief's Branch, Integrated HQ of MoD (Army), New Delhi | Shri Manoj Bapna Shri Ajay Kumar Sinha (Alternate) |
| MECON Limited, Ranchi | Shri Shankar Ray Shri Ayush Srivastava (Alternate) |
| Ministry of Ports, Shipping and Waterways, New Delhi | Shri H. N. Aswath Shri Anil Pruthi (Alternate) |
| Mumbai Port Trust, Mumbai | Dy Chief Engineer (Design) Superintending Engineer (Design) (Alternate) |
| Nagadi Consultants Pvt Limited, New Delhi | Dr V. V. S. Rao Shri N. Santosh Rao (Alternate) |
| National Capital Region Transport Corporation, New Delhi | Shri Jitender Kumar |
| National High Speed Rail Corporation Ltd, Mumbai | Representative |
| National Institute of Disaster Management, New Delhi | Dr Chandan Ghosh Dr Amir Ali Khan (Alternate) |
| NTPC Limited, Noida | Shri Mohit Jhalani |
| Power Grid Corporation of India Limited, Gurugram | Representative |
| Research Designs and Standards Organization (Ministry of Railways), Lucknow | Shri Sameer Singh Shri S. K. Ojha (Alternate) |
| RITES Limited, Gurugram | Shri Koshy Vaidyan Shri Sumeet Mahajan (Alternate) |
| Safe Enterprises, Mumbai | Shri Vikram Singh Rao Shri Suryaveer Singh Rao (Alternate) |
| STUP Consultants Pvt Ltd, Mumbai | Shri Anirban Sengupta Shri Yogesh Waingankar (Alternate) |
| Tata Consulting Engineers Limited, Mumbai | Shri Sanjeev Gupta Shri B. N. Nagaraj (Alternate) |
| Telangana State Research Laboratories, Hyderabad | Shri A. G. Manoj Kumar Shri Ashirwadam Jakkula (Alternate-I) Smt M. Manjula (Alternate-II) |
| The Pressure Piling Co (I) Pvt Limited, Mumbai | Shri V. C. Deshpande Shri Pushkar V. Deshpande (Alternate) |
| Unique Geocivil Services Pvt Ltd, Surat | Shri Nehal H. Desai Shri Hitesh H. Desai (Alternate-I) Shri Dhruval D. Shah (Alternate-II) |
| In Personal Capacity, 1-B, Villakkupattam Palace, First Floor, 48, New Avadi Road, Kilpauk, Chennai 600010 | Dr V. Balakumar |
| BIS Directorate General | Shri Dwaipayan Bhadra, Scientist ‘E’/ Director and Head (Civil Engineering) [Representing Director General (Ex-officio)] |
| Member Secretary Shri Dheeraj Damachya Scientist ‘B’ / Assistant Director (Civil Engineering), BIS |  |