**IS 7756 : 2024**

***भारतीय मानक***

***Indian Standard***

**ईएनटी शल्य चिकित्सा उपकरण — जीभ डिप्रेसर — लैक पैटर्न —**

**विशिष्टि**

*(*पहला पुनरीक्षण)

**ENT Surgery Instruments –– Tongue Depressor — Lack’s Pattern — Specification**

) *First Revision)*

ICS 11.040.30

 © BIS 2024

भारतीय मानक ब्यूरो

BUREAU OF INDIAN STANDARDS

मानक भवन, 9 बहादुर शाह ज़फर मार्ग, नई दिल्ली - 110002

MANAK BHAVAN, 9 BAHADUR SHAH ZAFAR MARG

NEW DELHI - 110002

[www.bis.gov.in](http://www.bis.org.in) [www.standardsbis.in](http://www.standardsbis.in)

 **XXXXXX Price Group X**

Ear, Nose, Throat and Head & Neck Surgery (ENT - H&N) Instruments Sectional Committee, MHD 04.

FOREWORD

This Indian Standard (Second Revision) was adopted by the Bureau of Indian Standards after the draft Finalized by the Ear, Nose, Throat and Head & Neck Surgery (ENT - H&N) Instruments Sectional Committee had been approved by the Medical Equipment and Hospital Planning Division Council.

This standard was originally published in 1975. The First revision of this standard has been brought out to align it with recent developments and to bring the standard in line with the latest style and format of Indian Standards.

The composition of the committee responsible for formulation of this standard is given in **Annexure- A**

For the purpose of deciding whether a particular requirement of this standard is complied with the final value, observed or calculated, expressing the result of a test or analysis shall be rounded off in accordance with IS 2: 2022 ‘Rules for rounding off numerical values (*second revision*)’. The number of significant places retained in the rounded-off value should be same as that of the specified value in this standard.

*Indian Standard*

ENT SURGERY INSTRUMENTS –– TONGUE DEPRESSOR — LACK’S PATTERN — SPECIFICATION

## SCOPE

This Standard cover dimensional and other requirements for Lack’s pattern’ tongue depressor.

## REFERENCES

The standards given below contain provisions, which, through reference in this text, constitute provisions of this standard. At the time of publication, the editions indicated were valid. All standards are subject to revision, and parties to agreements based on this standard are encouraged to investigate the possibility of applying the most recent editions of these standards.

|  |  |
| --- | --- |
| ***IS No.*** | ***Title*** |
| IS 7531: 1990 | Surgical instruments – Corrosion resistance of stainless steelsurgical instruments – Methods of tests *(first revision)* |
| IS 6911: 2017 | Stainless steel plate, sheet and strip – Specification *(second revision)* |

## SHAPE AND DIMENSIONS

* 1. The shape and dimensions of tongue depressor lack's pattern shall be as shown in **Fig. 1**.
	2. Deviation of ± 25 percent shall be allowed on all dimensions.

## MATERIAL

The material shall be stainless steel conforming to designation X04Cr18Ni9 or X07Crl8Ni9 of IS 6911.

## WORKMANSHIP AND FINISH

* 1. The depressor shall be of one piece and shall be free from burrs, pits, cracks and other surface flaws.
	2. The edges of the depressor shall be even, rounded and shall not be sharp.
	3. The depressor shall be polished bright and passivated.
	4. The thickness of the depressor shall be uniform throughout.

*W*

*W*1

*L*2

R25

2.0

80°

*L*1

R20

10

All dimensions in millimeters.

**FIG. 1 DEPRESSOR, TONGUE, LACK'S PATTER**

## TESTS

* 1. **Flexibility Test**

Grip the whole length of the straight limb of the depressor in a vice. Apply a force to the free end of the depressor to deflect it by an amount shown in **Fig. 1.** The depressor shall not have acquired any new permanent set. The blade shall show no sign of damage after test.

* 1. **Proof Load Test**

Grip the depressor firmly in a vice close to the large blade, with the blade protruding and facing upwards. Suspend gradually a load of 180 N (18 kgf approximately) from the tip of the blade for 2 minutes. Similarly carry out the test on the small blade using a load of 100 N (10 kgf approximately). After the test, the depressor shall show no sign of damage.

* 1. **Corrosion Resistance**

The depressor shall show no sign of corrosion when tested in accordance with IS 7531.

## MARKING

**7.1** Each depressor shall be marked with the manufacturer’s name, initials or recognized trademark and the country of manufacture.

## BIS Certification Marking

The product(s) conforming to the requirements of this standard may be certified as per the conformity assessment schemes under the provisions of the *Bureau of Indian Standards Act, 2016* and the Rules and Regulations framed thereunder, and the product(s) may be marked with the Standard Mark.

## PACKING

The depressor shall be wrapped in moisture-proof paper or packed in polyethylene bags. The depressor may also be packed as agreed to between the purchaser and the supplier. On the package of the product batch Number, Lot Number, and Serial Number shall be mentioned.

**ANNEX A**

(*Foreword*)

**COMMITTEE COMPOSITION**

| ***Organization*** | ***Representative(S)*** |
| --- | --- |
| All India Institute of Medical Sciences, New Delhi | Dr. Rakesh Kumar *(Chairperson)* |
| ALPS International, New Delhi | Shri.Alok Narang,  |
| Shri. Karan Narang (*Alternate*) |
| All India Institute of Medical Sciences, New Delhi | Dr. Prem Sagar  |
| Dr. Arvind Kumar (*Alternate*) |
| All India Institute of Medical Sciences, Bhopal | Dr. Vikas Gupta  |
| Dr. Ganakalyan Behera (*Alternate*) |
| Association of Indian Medical Device Industry, New Delhi | Shri. Tarlochan Dev  |
| Shri. Ankur Bhargava (*Alternate* I*)* |
| Dr. C.S. Prasad (*Alternate* II) |
| Directorate General of Health Services, New Delhi Central | Ms. Dr. Pallika Kumar |
| Government Medical College & Hospital, Chandigarh | Dr. Surinder K Singhal  |
| Shri. Nitin Gupta (*Alternate*) |
| Happy Reliable Surgeries Private Limited, Bangalore | Shri. Hemant Savale  |
| Shri. Sanjeev Gautam (*Alternate*) |
| India Medtronic Private Limited, Gurugram | Smt. LATIKA VATS  |
| Shri. Saurabh Sable (*Alternate* I*)* |
| Shri. Sandeep Verma (*Alternate* II) |
| Indian Institute of Technology Kanpur, Kanpur | Dr. A. R. Harish  |
| Kalam Institute of Health Technology, Vishakhapatnam | Dr. Arjun Thimmaiah  |
| Shri. Amit Sharma (*Alternate*) |
| Karl Storz Endoscopy India Private Limited, New Delhi | Shri. Sandeep Sethi |
| Shri. Kapil Rana (*Alternate*) |
| Serwell MediEquip, Chennai | Shri. T. Jebin Samuel  |
| Shri. R. Radhakrishnan (*Alternate* I*)* |
| Shri. G. Sathish (*Alternate* II) |
| Tata Memorial Center (Hospital), Mumbai | Dr. RICHA VAISH  |
| Shri. Vijay Yashwant Mestri (*Alternate* I*)* |
| Dr. Arjun Singh (*Alternate* II) |
| Postgraduate Institute of Medical Education and Research, Chandigarh | Smt. Dr Jaimanti Bakshi |
| In Personal Capacity , D-2 ,Tower 7, Type 5, East Kidwai Nagar, New Delhi-110023 | Dr. Kapil Sikka |
| BIS Directorate General | Shri A. R. Unnikrishnan Scientist ‘G’ And Head (Medical Equipment And Hospital Planning) [Representing Director General (Ex-Officio)] |

Ear, Nose, Throat and Head & Neck Surgery (ENT - H&N) Instruments Sectional

Committee, MHD 04

*Member Secretary*

Mr. Karthik Reddy Katipally

Scientist ‘B’/Assistant Director

(Medical Equipment And Hospital Planning). Bis