**Doc: PCD 01 (21671) F**

***भारतीय मानक***

***Indian Standard***

[**IS XXXXX : 2024**](https://www.services.bis.gov.in/php/BIS_2.0/StandardsFormulationV2/Upload3.php?ID=UUhaazE3ZldpOHl6TG9tS3dnbndDdz09)

**DME मिश्रित द्रवित पेट्रोलियम गैस (LPG)**

**ईंधन — गैस क्रोमैटोग्राफी द्वारा डाइमिथाइल ईथर**

**(DME) का मात्रात्मक निर्धारण**

**DME- Blended Liquified**

**Petroleum Gas (LPG) Fuel —**

**Quantitative Determination of**

**Dimethyl Ether (DME) Content**

**by Gas Chromatography**

ICS 75.160.30

© BIS 2024

भारतीय मानक ब्यूरो

BUREAU OF INDIAN STANDARDS

मानक भवन, 9 बहादुर शाह ज़फर मार्ग, नई दिल्ली - 110002

MANAK BHAVAN, 9 BAHADUR SHAH ZAFAR MARG

NEW DELHI - 110002

[www.bis.gov.in](http://www.bis.org.in) [www.standardsbis.in](http://www.standardsbis.in)

**July 2024 Price Group X**

Methods of Sampling and Test for Petroleum and Related Products of Natural or Synthetic Origin (excluding bitumen) Sectional Committee, PCD 01

FOREWORD

This Indian Standard was adopted by the Bureau of Indian Standards, after the draft finalized by the Methods of Sampling and Test for Petroleum and Related Products of Natural or Synthetic Origin (excluding bitumen) Sectional Committee had been approved by the Petroleum, Coal and Related Products Division Council.

Liquified petroleum gas (LPG) is being used in India to cater to the energy needs of domestic, commercial and industrial sectors apart from use as automotive fuel. The consumption of LPG is ever increasing in the country. This necessitates use of alternate fuels to partially substitute LPG with fuel such as Dimethyl ether (DME). It is the simplest ether with oxygen connecting two methyl groups having no C-C bond. DME can be blended with LPG up to 20 percent by weight and the blended fuel can be used for cooking in households and other applications.

As there is no Indian Standard available for determination of Dimethyl ether (DME) content in DME – Liquified petroleum gas (LPG) blended fuel, this standard has been formulated.

The composition of the Committee, Subcommittee responsible for the formulation of this standard is given in Annex B.

In reporting the result of a test or analysis made in accordance with this standard, if the final value, observed or calculated, is to be rounded off, it shall be done in accordance with IS 2 : 2022 ‘Rules for rounding off numerical values (*second revision*)’.

*Indian Standard*

DME- BLENDED LIQUIFIED PETROLEUM GAS (LPG) FUEL — QUANTITATIVE DETERMINATION OF DIMETHYL ETHER (DME) CONTENT BY GAS CHROMATOGRAPHY

**1 SCOPE**

This standard prescribes a method of test for quantitative determination of dimethyl ether (DME) content in DME-liquefied petroleum gas (LPG) blended fuel by gas chromatography. Component concentrations are determined in the range of 10 mass percent to 100 mass percent.

**2 REFERENCES**

The standards listed in Annex A contain provisions which, through reference in this text, constitute provisions of this standard. At the time of publication, the editions indicated were valid. All standards are subjected to revision and parties to agreements based on this standard are advised to use the latest edition of these standards:



**3 SUMMARY OF THE TEST METHOD**

**3.1 Outline of the method**

DME-LPG blended sample is analyzed via gas sampling valves by gas chromatography and compared to corresponding components separated under identical operating conditions from pure DME reference standard. The chromatogram of the sample is interpreted by comparing peak retention times and areas with those obtained for the pure DME reference standard.

**3.2 Apparatus**

**3.2.1** *Gas Chromatograph* (*GC*)

Any gas chromatographic instrument provided with a linear temperature programmable column oven. Multi-step column oven temperature programming is required, consisting of an initial hold time, an initial temperature program followed by an isothermal temperature hold and another programmed temperature rise. The temperature control shall be capable of obtaining a retention time repeatability of 0.05 min throughout the analysis. A flame ionization detector (FID) having sensitivity of 0.5 percent (mole) or less for the DME compound is strongly recommended.

**3.2.2** *Data Acquisition*

Any commercial integrator or computerized data acquisition system may be used for display of the chromatographic detector signal and peak area integration. The device shall be capable of calibration and reporting of the final response corrected results.

**3.2.3** *Sample Introduction*

For gas sampling, a six-port gas sampling valve (GSV) with a 250 μl fixed sampling loop may be provided. This valve shall be contained in a heated enclosure and operated at a temperature above the boiling point of the highest boiling component in the sample.

**3.3 Gas Controls**

The GC shall be provided with suitable facilities for delivery and control of carrier gas and the detector gases. This will consist of the appropriate tank and downstream regulators and supply tubing as well as the mass or pressure controls for the precise regulation of the instrument operation.

**3.4 Columns**

Condition all columns used according to the manufacturers constructions prior to use. The recommended analytical column for this test method is 100 m length, internal diameter 0.25 mm and film thickness 0.5 µm of 100 percent Dimethylpolysiloxane capillary column.

**4 REAGENTS AND MATERIALS**

**4.1 Carrier Gases**

For carrier gases, it is recommended to install commercial active oxygen scrubbers and water dryers, such as molecular sieves, ahead of the instrument to protect the system’s chromatographic columns. Follow manufacturer’s instructions in the use of such gas purifiers and replace as necessary.

**4.1.1** *Helium —* 99.995 percent minimum purity, < 0.1 ppm H2O. The use of appropriate scrubbers may be sufficient to obtain the desired purity.

**4.2 Detector Gases**

**4.2.1** *Hydrogen —* 99.995 percent minimum purity. The use of appropriate scrubbers may be sufficient to obtain the desired purity.

NOTE **—** Hydrogen is a flammable gas under high pressure.

**4.2.2** *Nitrogen —* 99.995 percent minimum purity. The use of appropriate scrubbers may be sufficient to obtain the desired purity.

**4.2.3** *Air —* less than 10 ppm each of total hydrocarbons and water. The use of appropriate scrubbers may be sufficient to obtain the desired purity.

NOTE **—** Improper handling of compressed gas cylinders containing air, nitrogen, hydrogen, or helium can result in an explosion. Rapid release of nitrogen or helium can result in asphyxiation.

**4.3 Reference Standards**

**4.3.1** *Purity of Reagents*

Reagent grade chemicals shall be used in all tests. Unless otherwise indicated, it is intended that all reagents conform to the specifications of the Committee on Analytical Reagents of the American Chemical Society where such specifications are available. Other grades may be used, provided it is first ascertained that the reagent is of sufficiently high purity to permit its use without lessening the accuracy of the determination.

**4.3.2** *Reference DME Standard*

Analytical grade DME reference standard commercially available and may be used to establish quantitative determination of DME content in DME-LPG blended fuel.

**5 PREPARATION OF APPARATUS**

**5.1** Set up the instrumentation in accordance with the manufacturer’s instructions or as specified herein.

**5.2** Install and condition the column according to manufacturer’s instructions.

**5.3** Set the GC instrument to the operating parameters. Allow the instrument to stabilize before proceeding with calibration and sample injections. Typical operating conditions for 100 percent dimethylpolysiloxane column are provided in Table 1.

**Table 1 Typical Operating Conditions**

(*Clauses* 5.3 *and* 7.3.1)

|  |  |
| --- | --- |
| Column | 100 m × 0.25 mm × 0.5 µm of 100 percent Dimethylpolysiloxane |
| Carrier Gas | Helium |
| Carrier Gas flow | 2.0 ml/min |
| Injector temperature | 250 °C |
| Injection type | Split with split ratio 150 : 1 |
| Purge flow | 3.0 ml/min |
| Injection volume | 250 µl |
| Oven program | Initial temperature 35 °C for 10 min;  First ramp at 2.5 °C/min to 120 °C hold for 0 min  Second ramp at 15 °C/min to 220 °C hold for 5 min |
| Detector | FID  Temperature: 250 °C Hydrogen flow: 40 ml/min  Air flow: 400 ml/min  Make-up gas: Nitrogen  Make-up gas flow: 30 ml/min |
| Analysis time | 55.6 min |

**5.4** Obtain duplicate chromatograms of the standard or sample, or both. Ensure that none of the peaks obtained have exceeded the upper range limit of the data handling device (at full scale on the data handling device, all peaks are on scale and display symmetrical, Gaussian shapes as opposed to flat peak tops). Use the same sample size (split ratio) and range for all runs. Example chromatograms are provided in Fig. 1.




Fig. 1 Example Chromatogram Using the Dimethylpolysiloxane Column

**5.5 Gas Sampling Valve**

Set valve on and off times to comply with manufacturer’s instructions

**6 CALIBRATION AND STANDARDIZATION**

**6.1 Qualitative**

Determine the retention time of DME by analyzing known reference standard in the same manner as the samples. Typical retention time of DME is 7.48 min.

**6.2 Quantitative**

Determine the quantity of DME in DME-LPG blended fuel is interpreted by comparing peak areas with those obtained for the pure DME reference standard.

1. **PROCEDURE**
   1. **Sampling**

Sampling at the sample source and at the chromatograph shall always be done in a manner that ensures that a representative sample is being analyzed. Lack of precision and accuracy in using this test method can most often be attributed to improper sampling procedures.

* 1. **Gas Sample Valve Injection**

Flush a gas sample loop with 5 ml to 10 ml of standard and sample (approximate 45s time), close cylinder valve, and allow the sample pressure to equilibrate to atmospheric pressure (stopped flow) before introducing the sample into the carrier gas stream.

* 1. **Sample Analysis Procedure**
     1. Adjust the instrument operating variables to the values specified in Table 1.
     2. Equilibrate the chromatographic system and inject the air blank until a representative chromatogram is obtained.
     3. Inject an appropriate size DME reference standard into the injection port and start the analysis. Obtain a chromatogram and a peak integration report. Repeat the same standard for six times consecutively.
     4. Inject a minimum of one air blank to check for carry over after six DME reference standards.
     5. Inject an appropriate size sample into the injection port and start the analysis. Obtain a chromatogram and a peak integration report. Repeat the sample in duplicate.
     6. Inject DME standard throughout the analysis, with a minimum of one injection at the end of sequence. One injection after every sixth sample analysis is recommended.
     7. Record the peak area response for each analyte of interest, if present. Note all observed peaks in the blank injection and exclude these as artifacts from any calculation.

1. **SYSTEM SUITABILITY**
   1. The percent RSD for peak area response of DME from the first six injections of the DME standard shall be ≤ 10 percent.
   2. The percent RSD for peak area response of DME from all injections of the DME standard throughout the analysis shall be ≤ 10 percent.
2. **CALCULATION**

**9.1** Identify DME peak by matching retention times with those for known reference standards. Obtain the area for DME peak.

DME content (percent *m/m*) = (*Asample* /*Astandard*) × purity of standard

where

*Asample* ***=*** Peak area of DME in sample; and

*Astandard* = Mean peak area of DME in first six standard injections.

**10 REPORT**

Report the concentration of DME component as percent (*m/m*), to the nearest 0.1 percent (*m/m*).

**11 PRECISION**

**11.1 Repeatability**

The difference in two test results obtained by the same operator with the same apparatus in a given laboratory under constant operating conditions on test samples taken from the same laboratory sample should, in the long run in the normal and correct operation of the test method not exceed the values given in Table 2.

**Table 2 Repeatability Values of the Method Obtained with the 20 Percent DME-80 Percent LPG Blended Fuel**

(*Clause* 11.1)

|  |  |  |  |
| --- | --- | --- | --- |
| **SI No.** | **Component** | **Concentration**  (percent *m/m*) | **Repeatability** |
| (1) | (2) | (3) | (4) |
| i) | DME | 18.5 | 0.7 |
| NOTE — Calculated average value of 6 replications. | | | |

**ANNEX A**

(*Clause* 2)

**LIST OF REFERRED STANDARDS**

|  |  |
| --- | --- |
|  |  |
|  | pgs |
|  | : |
|  | m |
|  |  |
|  |  |
|  |  |
|  | l— |
|  |  |
|  | d |

**ANNEX B**

(*Foreword*)

**COMMITTEE COMPOSITION**

Methods of Sampling and Test for Petroleum and Related Products of Natural or Synthetic Origin (excluding bitumen) Sectional Committee, PCD 01

| *Organization* |  | *Representative*(*s*) |
| --- | --- | --- |
| CSIR - Indian Institute of Petroleum, Dehradun |  | Dr Harender Singh Bisht **(*Chairperson*)** |
| Air Headquarters, Ministry of Defence, New Delhi |  | Wg Cdr Y. Bhardwaj  Wg Cdr VS Choudhary (*Alternate* I) |
| Bharat Petroleum Corporation Limited, Mumbai |  | Shri R. Subramanian  Shri C. Shanmuganathan *(Alternate*) |
| Castrol India Limited, Mumbai |  | Shri Raman Rai |
| Central Institute of Plastics Engineering and Technology, Bhubaneshwar |  | Dr Smita Mohanty  Dr R. Ananthakumar *(Alternate*) |
| Central Revenue Control Laboratory, New Delhi |  | Shri V. Suresh |
| Chennai Petroleum Corporation Limited, Chennai |  | Shri H. Ramakrishnan  Shri M. Balaguru (*Alternate*) |
| CSIR - Central Institute for Mining and Fuel Research, Dhanbad |  | Shri S. R. K. Rao  Shri P. K. Singh (*Alternate* I)  Shri S. Dutta (*Alternate* II) |
| CSIR - Indian Institute of Petroleum, Dehradun |  | Dr Pankaj Kumar Kanaujia  DR G. D. Thakre *(Alternate*) |
| Directorate General of Aeronautical Quality Assurance, Ministry of Defence, New Delhi |  | Shri Pankaj Chawla  Dr Mrinmoy Garai (*Alternate*) |
| Directorate General of Quality Assurance, Ministry of Defence, Kanpur |  | Dr Om Prakash Singh  Shri A. K. Kanaujia (*Alternate*) |
| Elico Limited, Hyderabad |  | Shri T. V. Shiva K. Rao  Shri N. Raju (*Alternate*) |
| GAIL (India) Limited, New Delhi |  | Dr Nityananda Panda  Dr Gopal Dayal (*Alternate*) |
| Gulf Oil Lubricants India Limited, Mumbai |  | Shri CT Chidambaram  Shri Mayuresh Godbole (*Alternate* I)  Shri S. Ganesh (*Alternate* II) |
| HPCL Mittal Energy Limited, Noida |  | DR Hemant Tyagi  Shri Narendra Kumar Gupta (*Alternate*) |
| Hindustan Petroleum Corporation Limited, Mumbai |  | Shri Elecheran Kumar  Shri Santosh Dhaku Bhogale *(Alternate* I)  Shri Mahesh Totla (*Alternate* II) |
| Indian National Ship-Owners Association, Mumbai |  | Shri Chitta Ranjan Dash  Shri Shrikant Shyamkant Madiwale (*Alternate*) |
| Indian Oil Corporation Limited — Refineries and Pipelines Division, New Delhi |  | Dr Ashutosh Mishra  Dr Shashi Pal Singh (*Alternate*) |
| Indian Oil Corporation (MKTG), Mumbai |  | Shri As Krishnamoorthy  Shri Sreekumar N. Veedu (*Alternate*) |
| Indian Oil Corporation (R and D Centre), Faridabad |  | Dr Ajay Kumar Arora |
| Lubrizol India Limited, Mumbai |  | Shri Anil Mane  Shrimati Reena Kuril (*Alternate*) |
| Mangalore Refinery and Petrochemical Limited, Mangaluru |  | Shri Yogeesha  Shri Anitha Shetty (*Alternate* I)  Shri R. M. Praksah (*Alternate* II) |
| Research Designs & Standards Organisation (RDSO), Lucknow |  | Shri Rajesh Srivastava  Shri Sonam Gupta (*Alternate*) |
| National Test House, Kolkata |  | Shri Vinod Kumar AmirchandRam  Ms Ishita Sur (*Alternate*) |
| Nayara Energy Limited, Mumbai |  | Shri Narhar Deshpande  Shri Ketankumar Patel (*Alternate*) |
| Numaligarh Refinery Limited, Golaghat |  | Shri K. Srinivas  Shri Partha Jyoti Sharma (*Alternate*) |
| Oil and Natural Gas Corporation Limited, New Delhi |  | Shri Gour Mohan Dass  Shrimati Leena John (*Alternate* I)  Shri Dinesh S. R. Reddy Kakuturi (*Alternate* II) |
| Oil India Limited, Duliajan |  | Shri Surajit Bora |
| Reliance Industries Limited, Mumbai |  | Shri Pramod Mall |
| Shriram Institute for Industrial Research, Delhi |  | Dr Mukesh Garg  Dr Pravesh Kumar (*Alternate*) |
| BIS Directorate General |  | Shrimati Meenal Passi Scientist ‘F’/Senior Director and Head (Petroleum, Coal and Related Product) [Representing Director General (*Ex-officio*)] |
| *Member Secretary*  Shri Hari Mohan Meena  Scientist ‘C’/Deputy Director  (Petroleum, Coal and Related Products), BIS | | |

Methods of Test for Gaseous Fuels Subcommittee, PCD 1:5

|  |  |
| --- | --- |
| *Organization* | *Representative*(*s*) |
| In Personal Capacity, New Delhi | Dr Anurag A. Gupta (***Convenor***) |
| CSIR - Indian Institute of Petroleum, Dehradun | Dr Sunil Kumar Pathak  Dr Ganesh Naik (*Alternate*) |
| GAIL (India) Limited, New Delhi | Dr Gopal Dayal |
| Hindustan Petroleum Corporation Limited, Mumbai | Shri Santosh Dhaku Bhogale |
| Indian Oil Corporation Limited, New Delhi | Shri M. Sithanathan |
| Mangalore Refinery and Petro Chemical Limited, Mangaluru | Shri Sudheer Pai M.  Shri Uday. B (*Alternate*) |
| Nayara Energy Limited, Mumbai | Shri Narhar Deshpande  Shri Arpan Shah (*Alternate*) |
| Oil and Natural Gas Corporation Limited, New Delhi | Shri Dinesh S. R. Reddy Kakuturi |
| Reliance India Limited, Mumbai | Shri Shrikant Shingte |
| Shell India Markets Private Limited, Mumbai | Shri Siva Kasturi |